PLASTIC FM1 Efficiency* Calibration Document

Version 1

29 September, 2006

*Please see PLASTIC FM2 Efficiency Calibration Document for descriptions of the efficiency rates

Table of Contents

FM1 SWS Protons

Efficiency Rate Before and After Threshold Change ……………………...2

Adjusted Rates………….……………………………………………….….4

Table of Proton Files…………………………………………………….....9

FM1 SWS Helium Ions…………………………………………………………....14

FM1 SWS Heavy Ions………………………………………………………….....21

FM1 SWS Summary………………………………………………………………37

FM1 Quadrants 1, 2, and 3………………………………………………………...39

FM1

The procedure for determining the total efficiency for FM1 is generally the same as for FM2.  However, the SF and SFR thresholds were decreased following Bern testing.  This means some additional scaling of the efficiencies (obtained at Bern) is required.  There is no frequency information for the UNH test beams, so while stop, single position, stop * single position , and triple coincidence efficiencies can be compared between the test locations, start efficiencies can not.

The threshold setting at Bern was high, so a high MCP voltage was necessary to get appreciable counts (above the threshold).  This is similar to operating at a low MCP voltage with a low threshold.  In other words, more particles exceed the threshold as MCP voltage increases.  Assuming the relationship between valid counts and change in MCP voltage is similar no matter what the threshold value, the Bern data that requires valid events can simply be shifted to lower MCP settings until it matches up with the data taken after the threshold change.  For example, if the Bern stop efficiency curves are shifted (to the left) by 400 V they are in good agreement with the post-Bern data.  Unlike FM2, where the micro-channel plates changed over time, the change in threshold should not affect the position rates (S_NO_POS and S_MULT_POS) because they do not require both SF and SFR to be counted.  Shifting the FM1 single position efficiency curve does not make sense.  Instead, we simply assume that in our region of interest the single position efficiency is equal to 1 (based on the post-threshold change single position efficiency curves).  The total efficiency curves can now be recalculated., with the assumption that post-Bern start efficiencies follow the same trends as the values from Bern (with the MCP shift).

*10 keV Bern H+ beam frequency is estimated.

Table of FM1 H+ Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM1-041221-1159.log | H+ | 10 | 20 | 3060 | 1faa | 0.7 (estimate) | -8 | 0 |
| Bern | PLA-FM1-041220-1918.log | H+ | 2 | 20 | 2870 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1900.log | H+ | 2 | 20 | 2911 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1853.log | H+ | 2 | 20 | 2941 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1848.log | H+ | 2 | 20 | 2983 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1845.log | H+ | 2 | 20 | 3020 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1836.log | H+ | 2 | 20 | 3053 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1843.log | H+ | 2 | 20 | 3060 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1841.log | H+ | 2 | 20 | 3095 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041220-1842.log | H+ | 2 | 20 | 3138 | 7F0E | 0.75 | -8 | 0 |
| Bern | PLA-FM1-041221-1128.log | H+ | 20 | 20 | 2906 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-1124.log | H+ | 20 | 20 | 2949 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-1121.log | H+ | 20 | 20 | 2988 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-1120.log | H+ | 20 | 20 | 3025 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-1113.log | H+ | 20 | 20 | 3060 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-1118.log | H+ | 20 | 20 | 3060 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-1116.log | H+ | 20 | 20 | 3099 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-1117.log | H+ | 20 | 20 | 3130 | 2FD7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0831.log | H+ | 60 | 20 | 2800 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0813.log | H+ | 60 | 20 | 2834 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0803.log | H+ | 60 | 20 | 2875 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0759.log | H+ | 60 | 20 | 2911 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0756.log | H+ | 60 | 20 | 2952 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0755.log | H+ | 60 | 20 | 2988 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0753.log | H+ | 60 | 20 | 3020 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0734.log | H+ | 60 | 20 | 3060 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0752.log | H+ | 60 | 20 | 3060 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0735.log | H+ | 60 | 20 | 3099 | 8FCB | 1 | -8 | 0 |
| Bern | PLA-FM1-041221-0738.log | H+ | 60 | 20 | 3129 | 8FCB | 1 | -8 | 0 |
| UNH | PLA-FM1-050418-2032.log | H+ (magnet gun) | 2 | 15 | 2616 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050418-2034.log | H+ (magnet gun) | 2 | 15 | 2679 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050418-2026.log | H+ (magnet gun) | 2 | 15 | 2758 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050418-2022.log | H+ (magnet gun) | 2 | 15 | 2882 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050418-2012.log | H+ (magnet gun) | 2 | 15 | 3000 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050418-2014.log | H+ (magnet gun) | 2 | 15 | 3083 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050418-2019.log | H+ (magnet gun) | 2 | 15 | 3150 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050418-2021.log | H+ (magnet gun) | 2 | 15 | 3220 | 7F0E |  | 21.4 |  |
| UNH | PLA-FM1-050601-0411.log | H+ (magnet gun) | 20 | 18 | 2586 |  |  | 6 |  |
| UNH | PLA-FM1-050601-0414.log | H+ (magnet gun) | 20 | 18 | 2657 |  |  | 6 |  |
| UNH | PLA-FM1-050601-0410.log | H+ (magnet gun) | 20 | 18 | 2721 |  |  | 6 |  |
| UNH | PLA-FM1-050601-0409.log | H+ (magnet gun) | 20 | 18 | 2818 |  |  | 6 |  |
| UNH | PLA-FM1-050601-0408.log | H+ (magnet gun) | 20 | 18 | 2909 |  |  | 6 |  |
| UNH | PLA-FM1-050601-0406.log | H+ (magnet gun) | 20 | 18 | 3016 |  |  | 6 |  |
| UNH | PLA-FM1-050601-0405.log | H+ (magnet gun) | 20 | 18 | 3068 |  |  | 6 |  |
| UNH | PLA-FM1-050601-0402.log | H+ (magnet gun) | 20 | 18 | 3209 |  |  | 6 |  |

We assume the threshold change affects all the species in a similar manner, except that for the heavier ions the single position efficiency approaches 0.9 rather than 1.  We base this on air beam scans at UNH after the threshold decrease.

Table of FM1 UNH Air Beam Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UNH | PLA-FM1-050411-1936.log | Air | 2 | 15 | 2500 | 7F0E |  |  |  |
| UNH | PLA-FM1-050411-1932.log | Air | 2 | 15 | 2610 | 7F0E |  | -15 |  |
| UNH | PLA-FM1-050411-1930.log | Air | 2 | 15 | 2796 | 7F0E |  | -15 |  |
| UNH | PLA-FM1-050411-1913.log | Air | 2 | 15 | 2873 | 7F0E |  | -15 |  |
| UNH | PLA-FM1-050411-1916.log | Air | 2 | 15 | 2977 | 7F0E |  | -15 |  |
| UNH | PLA-FM1-050411-1919.log | Air | 2 | 15 | 3063 | 7F0E |  | -15 |  |
| UNH | PLA-FM1-050411-1922.log | Air | 2 | 15 | 3150 | 7F0E |  | -15 |  |
| UNH | PLA-FM1-050411-1929.log | Air | 2 | 15 | 3209 | 7F0E |  | -15 |  |
| UNH | PLA-FM1-050411-1846.log | Air | 10 | 15 | 2603 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1842.log | Air | 10 | 15 | 2713 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1840.log | Air | 10 | 15 | 2785 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1820.log | Air | 10 | 15 | 2874 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1828.log | Air | 10 | 15 | 2981 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1835.log | Air | 10 | 15 | 3046 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1836.log | Air | 10 | 15 | 3141 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1838.log | Air | 10 | 15 | 3213 | 7F2B |  | -6.5 |  |
| UNH | PLA-FM1-050411-1800.log | Air | 20 | 15 | 2724 | 7F4F |  | -6.5 |  |
| UNH | PLA-FM1-050411-1722.log | Air | 20 | 15 | 2785 | 7F4F |  | -6.5 |  |
| UNH | PLA-FM1-050411-1733.log | Air | 20 | 15 | 2873 | 7F4F |  | -6.5 |  |
| UNH | PLA-FM1-050411-1738.log | Air | 20 | 15 | 2972 | 7F4F |  | -6.5 |  |
| UNH | PLA-FM1-050411-1744.log | Air | 20 | 15 | 3060 | 7F4F |  | -6.5 |  |
| UNH | PLA-FM1-050411-1749.log | Air | 20 | 15 | 3133 | 7F4F |  | -6.5 |  |
| UNH | PLA-FM1-050411-1755.log | Air | 20 | 15 | 3217 | 7F4F |  | -6.5 |  |
| UNH | PLA-FM1-050411-1856.log | Air | 35 | 15 | 2605 | 7F86 |  | -6.5 |  |
| UNH | PLA-FM1-050411-1851.log | Air | 35 | 15 | 2713 | 7F86 |  | -6.5 |  |
| UNH | PLA-FM1-050411-1636.log | Air | 35 | 15 | 2777 | 7F86 |  | -6.5 |  |
| UNH | PLA-FM1-050411-1620.log | Air | 35 | 15 | 3060 | 7F86 |  | -6.5 |  |
| UNH | PLA-FM1-050411-1626.log | Air | 35 | 15 | 3135 | 7F86 |  | -6.5 |  |
| UNH | PLA-FM1-050411-1630.log | Air | 35 | 15 | 3208 | 7F86 |  | -6.5 |  |

FM1:  He+

Table of FM1 He+ Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM1-041220-0959.log | He2+ | 60 | 20 | 2760 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0945.log | He2+ | 60 | 20 | 2800 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0939.log | He2+ | 60 | 20 | 2830 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0937.log | He2+ | 60 | 20 | 2877 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0935.log | He2+ | 60 | 20 | 2905 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0934.log | He2+ | 60 | 20 | 2950 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0932.log | He2+ | 60 | 20 | 2985 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0931.log | He2+ | 60 | 20 | 3025 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0928.log | He2+ | 60 | 20 | 3060 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0925.log | He2+ | 60 | 20 | 3100 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0927.log | He2+ | 60 | 20 | 3138 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-1147.log | He2+ | 20 | 20 | 2800 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1140.log | He2+ | 20 | 20 | 2836 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1138.log | He2+ | 20 | 20 | 2874 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1136.log | He2+ | 20 | 20 | 2908 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1135.log | He2+ | 20 | 20 | 2953 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1134.log | He2+ | 20 | 20 | 2956 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1132.log | He2+ | 20 | 20 | 3023 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1108.log | He2+ | 20 | 20 | 3060 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1128.log | He2+ | 20 | 20 | 3095 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1129.log | He2+ | 20 | 20 | 3132 | 2fd7 | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1431.log | He2+ | 2 | 20 | 2800 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1427.log | He2+ | 2 | 20 | 2836 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1423.log | He2+ | 2 | 20 | 2911 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1422.log | He2+ | 2 | 20 | 2944 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1420.log | He2+ | 2 | 20 | 2986 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1419.log | He2+ | 2 | 20 | 3024 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1417.log | He2+ | 2 | 20 | 3053 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1357.log | He2+ | 2 | 20 | 3060 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1412.log | He2+ | 2 | 20 | 3060 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1413.log | He2+ | 2 | 20 | 3095 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-1415.log | He2+ | 2 | 20 | 3132 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041214-1052.log | He+ | 2 | 20 | 2720 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1038.log | He+ | 2 | 20 | 2765.5 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1030.log | He+ | 2 | 20 | 2805 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1027.log | He+ | 2 | 20 | 2832 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1025.log | He+ | 2 | 20 | 2876 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1024.log | He+ | 2 | 20 | 2904 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1019.log | He+ | 2 | 20 | 2942 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-0957.log | He+ | 2 | 20 | 2989 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-0954.log | He+ | 2 | 20 | 3017.4 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-0949.log | He+ | 2 | 20 | 3058 | 7f08 | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-0959.log | He+ | 2 | 20 | 3099.8 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1002.log | He+ | 2 | 20 | 3131.4 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1004.log | He+ | 2 | 20 | 3172.8 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1007.log | He+ | 2 | 20 | 3210 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041214-1012.log | He+ | 2 | 20 | 3242.1 | 7f0e | 1.4 | -8 | 0 |
| Bern | PLA-FM1-041219-1849.log | He+ | 2 | 15 | 2800 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1838.log | He+ | 2 | 15 | 2836 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1832.log | He+ | 2 | 15 | 2877 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1829.log | He+ | 2 | 15 | 2910 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1827.log | He+ | 2 | 15 | 2951 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1826.log | He+ | 2 | 15 | 2979 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1824.log | He+ | 2 | 15 | 3023 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1822.log | He+ | 2 | 15 | 3060 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1816.log | He+ | 2 | 15 | 3092 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1819.log | He+ | 2 | 15 | 3129 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041219-1821.log | He+ | 2 | 15 | 3175 | 7f0e | 1 | -8 | 0 |
| Bern | PLA-FM1-041220-0824.log | He+ | 60 | 20 | 2760 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0807.log | He+ | 60 | 20 | 2797 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0759.log | He+ | 60 | 20 | 2841 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0755.log | He+ | 60 | 20 | 2870 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0753.log | He+ | 60 | 20 | 2914 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0751.log | He+ | 60 | 20 | 2946 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0750.log | He+ | 60 | 20 | 2987 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0748.log | He+ | 60 | 20 | 3023 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0745.log | He+ | 60 | 20 | 3060 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0741.log | He+ | 60 | 20 | 3095 | 8fcb | 0.7 | -8 | 0 |
| Bern | PLA-FM1-041220-0744.log | He+ | 60 | 20 | 3128 | 8fcb | 0.7 | -8 | 0 |

FM1:  Heavy Ions

Table of FM1 Ar Ion Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM1-041217-1329.log | Ar+ | 20 | 20 | 2800 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1518.log | Ar+ | 20 | 20 | 2800 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1450.log | Ar+ | 20 | 20 | 2841 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1434.log | Ar+ | 20 | 20 | 2872 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1421.log | Ar+ | 20 | 20 | 2904 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1413.log | Ar+ | 20 | 20 | 2944 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1406.log | Ar+ | 20 | 20 | 2988 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1401.log | Ar+ | 20 | 20 | 3022 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1331.log | Ar+ | 20 | 20 | 3060 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1356.log | Ar+ | 20 | 20 | 3060 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1338.log | Ar+ | 20 | 20 | 3099 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1348.log | Ar+ | 20 | 20 | 3138 | 2fd7 | 0.48 | -8 | 0 |
| Bern | PLA-FM1-041217-1245.log | Ar+ | 60 | 20 | 2763 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1242.log | Ar+ | 60 | 20 | 2792 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1239.log | Ar+ | 60 | 20 | 2837 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1236.log | Ar+ | 60 | 20 | 2868 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1233.log | Ar+ | 60 | 20 | 2912 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1231.log | Ar+ | 60 | 20 | 2945 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1230.log | Ar+ | 60 | 20 | 2984 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1228.log | Ar+ | 60 | 20 | 3024 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1227.log | Ar+ | 60 | 20 | 3060 | 8fcb | 0.25 | -8 | 0 |
| Bern | PLA-FM1-041217-1823.log | Ar2+ | 20 | 20 | 2725 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1815.log | Ar2+ | 20 | 20 | 2760 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1810.log | Ar2+ | 20 | 20 | 2800 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1807.log | Ar2+ | 20 | 20 | 2835 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1805.log | Ar2+ | 20 | 20 | 2871 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1802.log | Ar2+ | 20 | 20 | 2907 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1801.log | Ar2+ | 20 | 20 | 2945 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1759.log | Ar2+ | 20 | 20 | 2978 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1758.log | Ar2+ | 20 | 20 | 3017 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1737.log | Ar2+ | 20 | 20 | 3060 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1756.log | Ar2+ | 20 | 20 | 3060 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1750.log | Ar2+ | 20 | 20 | 3098 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1752.log | Ar2+ | 20 | 20 | 3134 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1754.log | Ar2+ | 20 | 20 | 3176 | 2fd7 | 0.58 | -8 | 0 |
| Bern | PLA-FM1-041217-1721.log | Ar2+ | 60 | 20 | 2720 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1720.log | Ar2+ | 60 | 20 | 2765 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1717.log | Ar2+ | 60 | 20 | 2795 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1715.log | Ar2+ | 60 | 20 | 2832 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1714.log | Ar2+ | 60 | 20 | 2879 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1712.log | Ar2+ | 60 | 20 | 2914 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1710.log | Ar2+ | 60 | 20 | 2950 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1707.log | Ar2+ | 60 | 20 | 2985 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1705.log | Ar2+ | 60 | 20 | 3027 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1652.log | Ar2+ | 60 | 20 | 3060 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1703.log | Ar2+ | 60 | 20 | 3060 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1656.log | Ar2+ | 60 | 20 | 3098 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1659.log | Ar2+ | 60 | 20 | 3126 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-1701.log | Ar2+ | 60 | 20 | 3176 | 8fcb | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050110-1512.log | Ar3+ | 20 | 20 | 2800 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1515.log | Ar3+ | 20 | 20 | 2838 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1517.log | Ar3+ | 20 | 20 | 2873 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1523.log | Ar3+ | 20 | 20 | 2907 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1524.log | Ar3+ | 20 | 20 | 2945 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1526.log | Ar3+ | 20 | 20 | 2977 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1528.log | Ar3+ | 20 | 20 | 3016 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1530.log | Ar3+ | 20 | 20 | 3067 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1532.log | Ar3+ | 20 | 20 | 3095 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1534.log | Ar3+ | 20 | 20 | 3127 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1536.log | Ar3+ | 20 | 20 | 3169 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1545.log | Ar3+ | 20 | 20 | 3206 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1549.log | Ar3+ | 20 | 20 | 3244 | 5f6e | 0.25 | -8 | 0 |
| Bern | PLA-FM1-050110-1137.log | Ar3+ | 60 | 20 | 2726 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1139.log | Ar3+ | 60 | 20 | 2757 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1141.log | Ar3+ | 60 | 20 | 2800 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1143.log | Ar3+ | 60 | 20 | 2838 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1145.log | Ar3+ | 60 | 20 | 2873 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1147.log | Ar3+ | 60 | 20 | 2907 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1150.log | Ar3+ | 60 | 20 | 2945 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1134.log | Ar3+ | 60 | 20 | 2958 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1156.log | Ar3+ | 60 | 20 | 2977 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1204.log | Ar3+ | 60 | 20 | 3016 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1207.log | Ar3+ | 60 | 20 | 3061 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1219.log | Ar3+ | 60 | 20 | 3095 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1221.log | Ar3+ | 60 | 20 | 3127 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1225.log | Ar3+ | 60 | 20 | 3169 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1229.log | Ar3+ | 60 | 20 | 3206 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1234.log | Ar3+ | 60 | 20 | 3244 | FF76 | 1.1 | -8 | 0 |
| Bern | PLA-FM1-050110-1716.log | Ar4+ | 60 | 20 | 2727 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1740.log | Ar4+ | 60 | 20 | 2760 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1742.log | Ar4+ | 60 | 20 | 2800 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1744.log | Ar4+ | 60 | 20 | 2833 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1746.log | Ar4+ | 60 | 20 | 2867 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1748.log | Ar4+ | 60 | 20 | 2912 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1750.log | Ar4+ | 60 | 20 | 2939 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1753.log | Ar4+ | 60 | 20 | 2986 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1756.log | Ar4+ | 60 | 20 | 3024 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1711.log | Ar4+ | 60 | 20 | 3056 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1759.log | Ar4+ | 60 | 20 | 3060 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1801.log | Ar4+ | 60 | 20 | 3098 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1804.log | Ar4+ | 60 | 20 | 3127 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1809.log | Ar4+ | 60 | 20 | 3169 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1814.log | Ar4+ | 60 | 20 | 3206 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050110-1818.log | Ar4+ | 60 | 20 | 3244 | FF76 | 0.23 | -8 | 0 |
| Bern | PLA-FM1-050111-1044 | Ar5+ | 60 | 20 | 2684 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1041 | Ar5+ | 60 | 20 | 2725 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1040 | Ar5+ | 60 | 20 | 2760 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1038 | Ar5+ | 60 | 20 | 2800 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1035 | Ar5+ | 60 | 20 | 2833 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1031 | Ar5+ | 60 | 20 | 2867 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1030 | Ar5+ | 60 | 20 | 2912 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1027 | Ar5+ | 60 | 20 | 2939 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1025 | Ar5+ | 60 | 20 | 2986 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1020 | Ar5+ | 60 | 20 | 3060 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1007 | Ar5+ | 60 | 20 | 3098 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1010 | Ar5+ | 60 | 20 | 3127 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1012 | Ar5+ | 60 | 20 | 3173 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1014 | Ar5+ | 60 | 20 | 3210 | 8FCB | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050111-1017 | Ar5+ | 60 | 20 | 3242 | 8FCB | 0.2 | -8 | 0 |

Table of FM1 O Ion Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM1-041217-0823.log | O+ | 60 | 20 | 2687 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0820.log | O+ | 60 | 20 | 2723 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0751.log | O+ | 60 | 20 | 2754 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0818.log | O+ | 60 | 20 | 2760 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0816.log | O+ | 60 | 20 | 2803 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0814.log | O+ | 60 | 20 | 2836 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0813.log | O+ | 60 | 20 | 2878 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0811.log | O+ | 60 | 20 | 2909 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0808.log | O+ | 60 | 20 | 2950 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0806.log | O+ | 60 | 20 | 2978 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0804.log | O+ | 60 | 20 | 3024 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0802.log | O+ | 60 | 20 | 3055 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0755.log | O+ | 60 | 20 | 3060 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-0800.log | O+ | 60 | 20 | 3097 | 8fcb | 0.33 | -8 | 0 |
| Bern | PLA-FM1-041217-1000.log | O+ | 20 | 20 | 2756 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0948.log | O+ | 20 | 20 | 2800 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0942.log | O+ | 20 | 20 | 2837 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0937.log | O+ | 20 | 20 | 2870 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0934.log | O+ | 20 | 20 | 2903 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0931.log | O+ | 20 | 20 | 2940 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0928.log | O+ | 20 | 20 | 2978 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0924.log | O+ | 20 | 20 | 3018 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0907.log | O+ | 20 | 20 | 3060 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0921.log | O+ | 20 | 20 | 3061 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0912.log | O+ | 20 | 20 | 3088 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0915.log | O+ | 20 | 20 | 3136 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-041217-0919.log | O+ | 20 | 20 | 3164 | 2fd7 | 0.2 | -8 | 0 |
| Bern | PLA-FM1-050112-1319.log | O2+ | 60 | 20 | 3018 | 8FCB | 0.005 | -8 | 0 |
| Bern | PLA-FM1-050112-1236.log | O2+ | 60 | 20 | 3060 | 8FCB | 0.005 | -8 | 0 |
| Bern | PLA-FM1-050112-1244.log | O2+ | 60 | 20 | 3090 | 8FCB | 0.005 | -8 | 0 |
| Bern | PLA-FM1-050112-1250.log | O2+ | 60 | 20 | 3127 | 8FCB | 0.005 | -8 | 0 |
| Bern | PLA-FM1-050112-1255.log | O2+ | 60 | 20 | 3175 | 8FCB | 0.005 | -8 | 0 |
| Bern | PLA-FM1-050112-1259.log | O2+ | 60 | 20 | 3209 | 8FCB | 0.005 | -8 | 0 |
| Bern | PLA-FM1-050112-1305.log | O2+ | 60 | 20 | 3247 | 8FCB | 0.005 | -8 | 0 |

FM1 Solar Wind Sector Summary

FM1:  Other Quadrants

Table of FM1 Quadrant 1 Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM1-050118-1636.log | H+ | 3 | 20 | 3177 | 5F18 |  |  |  |
| Bern | PLA-FM1-050118-1614.log | H+ | 3 | 20 | 3251 | 5F18 |  |  |  |
| Bern | PLA-FM1-050118-1618.log | H+ | 3 | 20 | 3260 | 5F18 |  |  |  |
| Bern | PLA-FM1-050118-1624.log | H+ | 3 | 20 | 3290 | 5F18 |  |  |  |
| Bern | PLA-FM1-050118-1630.log | H+ | 3 | 20 | 3334 | 5F18 |  |  |  |
| Bern | PLA-FM1-050118-1610.log | H+ | 60 | 20 | 3147 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1717.log | H+ | 60 | 20 | 3183 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1658.log | H+ | 60 | 20 | 3227 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1611.log | H+ | 60 | 20 | 3251 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1706.log | H+ | 60 | 20 | 3258 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1725.log | H+ | 60 | 20 | 3260 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1730.log | H+ | 60 | 20 | 3260 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1731.log | H+ | 60 | 20 | 3260 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1733.log | H+ | 60 | 20 | 3260 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1735.log | H+ | 60 | 20 | 3260 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1710.log | H+ | 60 | 20 | 3300 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1714.log | H+ | 60 | 20 | 3348 | 7FE4 |  |  |  |
| Bern | PLA-FM1-050118-1723.log | H+ | 60 | 20 | 3363 | 7FE4 |  |  |  |
| UNH | PLA-FM1-050601-0415.log | H+ (magnet gun) | 20 | 18 | 2670 | 7FS0 |  | 59 |  |
| UNH | PLA-FM1-050601-0417.log | H+ (magnet gun) | 20 | 18 | 2761 | 7FS0 |  | 59 |  |
| UNH | PLA-FM1-050601-0419.log | H+ (magnet gun) | 20 | 18 | 3020 | 7FS0 |  | 59 |  |
| UNH | PLA-FM1-050601-0421.log | H+ (magnet gun) | 20 | 18 | 3124 | 7FS0 |  | 59 |  |
| UNH | PLA-FM1-050601-0423.log | H+ (magnet gun) | 20 | 18 | 3225 | 7FS0 |  | 59 |  |
| UNH | PLA-FM1-050601-0424.log | H+ (magnet gun) | 20 | 18 | 3320 | 7FS0 |  | 59 |  |

Quadrants 2 and 3 use the other TAC board, so they have a different correction (than quadrants 0 and 1) to the MCP voltage setting.

Table of FM1 Quadrant 2 Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM1-050115-2235.log | H+ | 3 | 20 | 2805 | SF18 |  |  |  |
| Bern | PLA-FM1-050115-2232.log | H+ | 3 | 20 | 2856 | SF18 |  |  |  |
| Bern | PLA-FM1-050115-2231.log | H+ | 3 | 20 | 2883 | SF18 |  |  |  |
| Bern | PLA-FM1-050115-2229.log | H+ | 3 | 20 | 2923 | SF18 |  |  |  |
| Bern | PLA-FM1-050115-2224.log | H+ | 3 | 20 | 2995 | SF18 |  |  |  |
| Bern | PLA-FM1-050115-2226.log | H+ | 3 | 20 | 3031 | SF18 |  |  |  |
| Bern | PLA-FM1-050115-2227.log | H+ | 3 | 20 | 3075 | SF18 |  |  |  |
| Bern | PLA-FM1-050115-1957.log | H+ | 60 | 20 | 2820 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1955.log | H+ | 60 | 20 | 2856 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1954.log | H+ | 60 | 20 | 2895 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1935.log | H+ | 60 | 20 | 2910 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1939.log | H+ | 60 | 20 | 2910 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1944.log | H+ | 60 | 20 | 2915 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1951.log | H+ | 60 | 20 | 2920 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-2008.log | H+ | 60 | 20 | 2930 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1946.log | H+ | 60 | 20 | 2954 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-2000.log | H+ | 60 | 20 | 2956 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-2006.log | H+ | 60 | 20 | 2956 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1947.log | H+ | 60 | 20 | 3004 | 8FCB |  |  |  |
| Bern | PLA-FM1-050115-1949.log | H+ | 60 | 20 | 3034 | 8FCB |  |  |  |
| UNH | PLA-FM1-050418-2149.log | H+ (magnet gun) | 2 | 15 | 2601 |  |  | 170 |  |
| UNH | PLA-FM1-050418-2151.log | H+ (magnet gun) | 2 | 15 | 2689 |  |  | 170 |  |
| UNH | PLA-FM1-050418-2147.log | H+ (magnet gun) | 2 | 15 | 2895 |  |  | 170 |  |
| UNH | PLA-FM1-050418-2145.log | H+ (magnet gun) | 2 | 15 | 3075 |  |  | 170 |  |
| UNH | PLA-FM1-050418-2141.log | H+ (magnet gun) | 2 | 15 | 3226 |  |  | 170 |  |

Table of FM1 Quadrant 3 Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM1-050121-1647.log | H+ | 3 | 20 | 2777 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1634.log | H+ | 3 | 20 | 2806 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1631.log | H+ | 3 | 20 | 2882 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1355.log | H+ | 3 | 20 | 2884 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1639.log | H+ | 3 | 20 | 2920 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1642.log | H+ | 3 | 20 | 2935 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1643.log | H+ | 3 | 20 | 2955 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1628.log | H+ | 3 | 20 | 2963 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1645.log | H+ | 3 | 20 | 2990 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1625.log | H+ | 3 | 20 | 3065 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1406.log | H+ | 3 | 20 | 3067 | 5F18 |  |  |  |
| Bern | PLA-FM1-050121-1341.log | H+ | 20 | 20 | 2774 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1336.log | H+ | 20 | 20 | 2809 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1332.log | H+ | 20 | 20 | 2839 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1326.log | H+ | 20 | 20 | 2880 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1353.log | H+ | 20 | 20 | 2884 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1328.log | H+ | 20 | 20 | 2925 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1350.log | H+ | 20 | 20 | 2925 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1331.log | H+ | 20 | 20 | 2937 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1324.log | H+ | 20 | 20 | 2975 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1322.log | H+ | 20 | 20 | 3070 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1318.log | H+ | 20 | 20 | 3152 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1301.log | H+ | 20 | 20 | 3225 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1312.log | H+ | 20 | 20 | 3258 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1315.log | H+ | 20 | 20 | 3300 | 2FD7 |  |  |  |
| Bern | PLA-FM1-050121-1034.log | H+ | 60 | 20 | 2728 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1024.log | H+ | 60 | 20 | 2770 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1020.log | H+ | 60 | 20 | 2811 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1018.log | H+ | 60 | 20 | 2852 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1049.log | H+ | 60 | 20 | 2880 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1047.log | H+ | 60 | 20 | 2886 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1015.log | H+ | 60 | 20 | 2890 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1045.log | H+ | 60 | 20 | 2914 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1014.log | H+ | 60 | 20 | 2960 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1012.log | H+ | 60 | 20 | 3000 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1010.log | H+ | 60 | 20 | 3030 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-0943.log | H+ | 60 | 20 | 3060 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-0946.log | H+ | 60 | 20 | 3065 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-0951.log | H+ | 60 | 20 | 3066 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1007.log | H+ | 60 | 20 | 3067 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-0956.log | H+ | 60 | 20 | 3110 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1000.log | H+ | 60 | 20 | 3143 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1002.log | H+ | 60 | 20 | 3176 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1003.log | H+ | 60 | 20 | 3226 | 8FCB |  |  |  |
| Bern | PLA-FM1-050121-1005.log | H+ | 60 | 20 | 3251 | 8FCB |  |  |  |
| UNH | PLA-FM1-050418-2129.log | H+ (magnet gun) | 2 | 15 | 2590 | 7F0E |  | -138 |  |
| UNH | PLA-FM1-050418-2140.log | H+ (magnet gun) | 2 | 15 | 2626 | 7F0E |  | -138 |  |
| UNH | PLA-FM1-050418-2126.log | H+ (magnet gun) | 2 | 15 | 2683 | 7F0E |  | -138 |  |