PLASTIC FM2 Efficiency Calibration Document

Version 1

29 September, 2006

Table of Contents

FM2 SWS Protons

Efficiency Rate Definitions………………………………………………...2

Changes Over Time…….……………………………………………….….5

Noise………………………….…………………………………………….8

Azimuth Response……………………………………………………….....8

Triple Coincidence………………………………………………………....9

Table of Proton Files……………………………………………………....11

FM2 SWS Helium Ions…………………………………………………………....13

FM2 SWS Heavy Ions………………………………………………………….....20

FM2 SWS Summary………………………………………………………………26

FM2 Quadrants 1, 2, and 3………………………………………………………...29

Efficiency Rate Definitions

There are three basic efficiencies that are combined to estimate the overall efficiency of ion detection by the PLASTIC instruments.

“Stop” Efficiency:  SFR/SF

The ratio of the start-stop coincidence rate (SFR) to the start rate (SF) is affected by the active area, gain, and voltage of the Micro-Channel Plates (MCPs); and by signal threshold and scattering.  In the plot below the FM2 Solar Wind Sector (SWS) stop efficiency is shown for protons at several beam energies and post-acceleration voltages (PACs).  The SFR0/SF0 ratio increases with MCP voltage until it levels off around 2900 V. (The “0” following the rate name indicates it is from TAC board 0.)  The 2 keV proton beam data was collected a few days after FM2 was installed in the vacuum chamber, and shows a stop efficiency higher than expected.  Had the 2 keV setting been repeated later, we expect the stop profile would nearly agree with the 1 keV proton beam curves.

FM2 SWS proton stop efficiency.  Data is from calibration testing at the University of Bern (Summer 2005).

“Start” Efficiency:  SF/(incident ion frequency) = SF/(active area * frequency/cm2)

A beam scanner was used to estimate the frequency per square centimeter of the incoming test beam.  The rate of ions incident on the TOF (Time of Flight) Section is obtained by multiplying the average frequency per square centimeter value by the active area of the entrance system (i.e. the electrostatic analyzer).   The active areas were measured to be 0.897 cm2 for the main channel of FM1, and 0.755 cm2 for FM2 (Reto Karrer, 2005).  The active area is different from the geometric area due to physical blockages, such as the support structure.  The incident ion rate includes this geometric effect.  Dividing the start frequency (SF) by the incident ion rate gives the start efficiency.  Similar to the stop efficiency plot above, the start efficiency plot shows an increase with MCP voltage and then a plateau.  The start rate levels out around 2700 V, which is less than the voltage required for maximum stop efficiency.  (Because beam scans were not available for all files in this document, start efficiency is not always included.)

FM2 SWS proton start efficiency. Data is from calibration testing at the University of Bern (Summer 2005).

“Single Position” Efficiency:  (S_E_NOT_REQ – S_NO_POS – S_MULT_POS)/SFR

The number of counts registered at a single position divided by the start-stop coincidence rate.  “S_E_NOT_REQ” indicates than an energy measurement is not required for an event to be counted, though it does require both a start (SF) and start-stop coincidence (SFR).  “S_NO_POS” is an event that registered without a position indication, while “S_MULT_POS” is an event that triggered multiple position responses.  (The position rates do not require both SF and SFR.)  The single position efficiency is particularly susceptible to changes in the MCP voltage, as cross-talk between the pixels will cause multiple position triggers for the same event.  As shown below the single position efficiency rises to nearly 100%, which is maintained over a range of several hundred volts, and then drops off as multiple position triggers increase at the higher MCP settings.

FM2 SWS proton single position efficiency. Data is from calibration testing at the University of Bern (Summer 2005).

“Total” Efficiency:   Stop * Start * Single Position

Taking the product of the above three efficiencies gives the total efficiency for an event recorded at a single position.

FM2 SWS proton single position total efficiency. Data is from calibration testing at the University of Bern (Summer 2005).

An optimal MCP voltage setting would maximize the three separate efficiencies, but such a setting does not necessarily exist.  Thus, a voltage is chosen that maximizes the total efficiency.  Once the MCP voltage has been selected, the total efficiency can be plotted against energy.  If a curve can be fit to the plot, then efficiency estimates for a range of energies can be obtained.  This procedure must be carried out for hydrogen (protons), light ions (such as helium), and heavier ions (such as argon and neon).

Changes Over Time

There is another factor that must be considered:  the response of the MCPs varies with time.  The plots above were made with data taken at the University of Bern in June 2005.  The efficiency curves for testing at UNH and Bern in 2005 show good agreement, but the curve for stop efficiency the following year at UNH shows a shift in instrument response.   For example, the efficiency for an MCP voltage of 2600 V at Bern in 2005 would require setting the MCP voltage to more than 3000 volts at UNH the following year during the final calibration tests.

Plots of stop efficiency from Bern (protons) and UNH (protons selected from air beam using a magnet) in 2005 show good agreement, but one year later the curve has shifted to the right by several hundred volts as a result of gradual changes over time in the MCPs.

Taking the changes over time into account, the best MCP voltage setting is at least 3100 volts.  No significantly higher voltages were tested.  Using the data taken at Bern with an MCP voltage setting of about 2650 volts gives the best estimate for the efficiency at the beginning of mission operations (with the actual MCP voltage set to 3100 V).

FM2 SWS proton stop efficiency curve for most likely MCP setting at the beginning of the mission.  The points near 20 kV that lie above the general trend are questionable because they were recorded a short time after the PLASTIC instrument was installed in the vacuum chamber.

FM2 SWS proton start efficiency curve for most likely MCP setting at the beginning of the mission. The points near 20 kV that lie above the general trend are questionable because they were recorded a short time after the PLASTIC instrument was installed in the vacuum chamber.

FM2 SWS proton single position efficiency curve for most likely MCP setting at the beginning of the mission.

FM2 SWS proton single position total efficiency curve for most likely MCP setting at the beginning of the mission.

Noise

How good is the efficiency estimate?  Plotting the time of flight spectra for the files used to generate the total efficiency plot, there is a small noise peak around channel 21.  This accounts for 1% or less of the total counts for each file.

FM2 SWS Time of Flight Spectra for protons at most likely MCP setting.

Azimuth Response

Examining the total efficiency across the solar wind sector shows a drop near the center of the section (azimuth angle = 0).  There is a spoke that supports the carbon foil at this location, and it scatters the incoming beam.  This results in a decrease in the single position efficiency.

FM2 SWS proton single position efficiency plotted over the possible range of azimuth angles.

Triple Coincidence

“Triple Coincidence” Efficiency:  S_E_REQ/SFR

Another rate of interest is the “triple coincidence efficiency.”  This is similar to the stop efficiency, except that in addition to SFR and SR it requires a single energy measurement.  Multiple energy events and events with no energy measurement are excluded.

FM2 SWS proton triple coincidence stop efficiency. Data is from calibration testing at the University of Bern (Summer 2005).

FM2 SWS proton triple coincidence efficiency plotted for the most likely MCP setting.

Table of FM2 H+ Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM2-050624-1654.log | H+ | 1 | 15 | 2460 | 06BC | 0.08 (0.015) | -8 | 0 |
| Bern | PLA-FM2-050624-1650.log | H+ | 1 | 15 | 2548 | 06BC | 0.08 (0.015) | -8 | 0 |
| Bern | PLA-FM2-050624-1658.log | H+ | 1 | 15 | 2644 | 06BC | 0.08 (0.015) | -8 | 0 |
| Bern | PLA-FM2-050624-1702.log | H+ | 1 | 15 | 2848 | 06BC | 0.08 (0.015) | -8 | 0 |
| Bern | PLA-FM2-050624-1704.log | H+ | 1 | 15 | 2990 | 06BC | 0.08 (0.015) | -8 | 0 |
| Bern | PLA-FM2-050624-1947.log | H+ | 1 | 18 | 2460 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-1941.log | H+ | 1 | 18 | 2549 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-1938.log | H+ | 1 | 18 | 2646 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-1936.log | H+ | 1 | 18 | 2756 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-1929.log | H+ | 1 | 18 | 2853 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-1927.log | H+ | 1 | 18 | 2990 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-2012.log | H+ | 1 | 20 | 2461 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-2006.log | H+ | 1 | 20 | 2555 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-2003.log | H+ | 1 | 20 | 2648 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-2002.log | H+ | 1 | 20 | 2760 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-2000.log | H+ | 1 | 20 | 2850 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050624-1958.log | H+ | 1 | 20 | 2990 | 06BC | 0.08 (0.05) | -8 | 0 |
| Bern | PLA-FM2-050618-0636.log | H+ | 2 | 18 | 2350 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0628.log | H+ | 2 | 18 | 2460 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0647.log | H+ | 2 | 18 | 2517 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0650.log | H+ | 2 | 18 | 2568 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0651.log | H+ | 2 | 18 | 2626 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0653.log | H+ | 2 | 18 | 2696 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0654.log | H+ | 2 | 18 | 2760 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0656.log | H+ | 2 | 18 | 2846 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0657.log | H+ | 2 | 18 | 2934 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050618-0658.log | H+ | 2 | 18 | 2992 | 7F0E | 0.32 | -8 |  |
| Bern | PLA-FM2-050617-1628.log | H+ | 2 | 18 | 2348 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1603.log | H+ | 2 | 18 | 2466 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1607.log | H+ | 2 | 18 | 2517 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1612.log | H+ | 2 | 18 | 2572 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1609.log | H+ | 2 | 18 | 2630 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1613.log | H+ | 2 | 18 | 2698 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1615.log | H+ | 2 | 18 | 2757 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1617.log | H+ | 2 | 18 | 2848 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050617-1627.log | H+ | 2 | 18 | 2955 | 7F0E | 0.38 | -8 |  |
| Bern | PLA-FM2-050624-1434.log | H+ | 60 | 20 | 2294 | scan | 0.2 | -8 | 0 |
| Bern | PLA-FM2-050624-1419.log | H+ | 60 | 20 | 2354 | scan | 0.2 | -8 | 0 |
| Bern | PLA-FM2-050624-1417.log | H+ | 60 | 20 | 2460 | scan | 0.2 | -8 | 0 |
| Bern | PLA-FM2-050624-1415.log | H+ | 60 | 20 | 2554 | scan | 0.2 | -8 | 0 |
| Bern | PLA-FM2-050624-1413.log | H+ | 60 | 20 | 2639 | scan | 0.2 | -8 | 0 |
| Bern | PLA-FM2-050624-1411.log | H+ | 60 | 20 | 2760 | scan | 0.2 | -8 | 0 |
| Bern | PLA-FM2-050624-1407.log | H+ | 60 | 20 | 2850 | scan | 0.2 | -8 | 0 |
| Bern | PLA-FM2-050624-1406.log | H+ | 60 | 20 | 2990 | scan | 0.2 | -8 | 0 |
| UNH | PLA-FM2-060515-2028.log | air (big ion gun) | 2 | 18 | 2760 | 7F0E |  | -2.5 |  |
| UNH | PLA-FM2-060515-2035.log | air (big ion gun) | 2 | 18 | 2915 | 7F0E |  | -2.5 |  |
| UNH | PLA-FM2-060515-2152.log | air (big ion gun) | 2 | 18 | 3000 | 7F0E |  | -2.5 |  |
| UNH | PLA-FM2-060515-1948.log | air (big ion gun) | 2 | 18 | 3024 | 7F0E |  | -2.5 |  |
| UNH | PLA-FM2-060515-2010.log | air (big ion gun) | 2 | 18 | 3140 | 7F0E |  | -2.5 |  |
| UNH | PLA-FM2-060515-2159.log | air (big ion gun) | 2 | 18 | 3146 | 7F0E |  | -2.5 |  |
| UNH | PLA-FM2-050607-0423.log | H+ (magnet gun) | 2 | 18 | 2273 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0424.log | H+ (magnet gun) | 2 | 18 | 2319 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0421.log | H+ (magnet gun) | 2 | 18 | 2361 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0419.log | H+ (magnet gun) | 2 | 18 | 2449 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0418.log | H+ (magnet gun) | 2 | 18 | 2550 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0416.log | H+ (magnet gun) | 2 | 18 | 2631 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0414.log | H+ (magnet gun) | 2 | 18 | 2722 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0412.log | H+ (magnet gun) | 2 | 18 | 2815 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0411.log | H+ (magnet gun) | 2 | 18 | 2909 | 7F0E |  | 19? |  |
| UNH | PLA-FM2-050607-0408.log | H+ (magnet gun) | 2 | 18 | 3001 | 7F0E |  | 19? |  |
| Bern | PLA-FM2-050617-1808.log | H+ | 2 | 18 | 2992 | 7F0E | 0.365 | scan |  |
| Bern | PLA-FM2-050624-1449.log | H+ | 60 | 18 | 3000 | 7FE8 | 0.2 | scan | 0 |

FM2:  Helium Ions

FM2 SWS helium ion stop efficiency.  Data is from calibration testing at the University of Bern (Summer 2005).

FM2 SWS helium ion start efficiency.  Data is from calibration testing at the University of Bern (Summer 2005).

FM2 SWS helium ion single position efficiency.  Data is from calibration testing at the University of Bern (Summer 2005).

FM2 SWS helium ion single position totoa efficiency.  Data is from calibration testing at the University of Bern (Summer 2005).

FM2 SWS helium stop efficiency curve for MCP setting equivalent to ~3100 V at UNH in Summer 2006.

FM2 SWS helium start efficiency curve for MCP setting equivalent to ~3100 V at UNH in Summer 2006.

FM2 SWS helium single position efficiency curve for MCP setting equivalent to ~3100 V at UNH in Summer 2006.

FM2 SWS helium single position total efficiency curve for MCP setting equivalent to ~3100 V at UNH in Summer 2006.

FM2 SWS helium triple coincidence stop efficiency. Data is from calibration testing at the University of Bern (Summer 2005).

FM2 SWS triple coincidence stop efficiency curve for MCP setting equivalent to ~3100 V at UNH in Summer 2006.

Table of FM2 Helium Ion Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM2-050618-1803.log | He+ | 60 | 18 | 2274 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1757.log | He+ | 60 | 18 | 2354 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1751.log | He+ | 60 | 18 | 2463 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1812.log | He+ | 60 | 18 | 2509 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1814.log | He+ | 60 | 18 | 2573 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1815.log | He+ | 60 | 18 | 2620 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1817.log | He+ | 60 | 18 | 2703 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1818.log | He+ | 60 | 18 | 2755 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1819.log | He+ | 60 | 18 | 2842 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1821.log | He+ | 60 | 18 | 2933 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050618-1823.log | He+ | 60 | 18 | 2988 | 7FE4 | 0.2 | -8 |  |
| Bern | PLA-FM2-050619-0627.log | He+ | 2 | 15 | 2354 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0619.log | He+ | 2 | 15 | 2456 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0613.log | He+ | 2 | 15 | 2517 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0608.log | He+ | 2 | 15 | 2572 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0606.log | He+ | 2 | 15 | 2630 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0604.log | He+ | 2 | 15 | 2700 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0602.log | He+ | 2 | 15 | 2762 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0601.log | He+ | 2 | 15 | 2851 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0554.log | He+ | 2 | 15 | 2944 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0545.log | He+ | 2 | 15 | 2995 | 7F0E | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0833.log | He+ | 1 | 20 | 2354 | 06BC | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0829.log | He+ | 1 | 20 | 2464 | 06BC | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0824.log | He+ | 1 | 20 | 2555 | 06BC | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0749.log | He+ | 1 | 20 | 2647 | 06BC | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0746.log | He+ | 1 | 20 | 2758 | 06BC | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0744.log | He+ | 1 | 20 | 2850 | 06BC | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050619-0743.log | He+ | 1 | 20 | 3000 | 06BC | 0.11 | -8 | 0 |
| Bern | PLA-FM2-050620-1044.log | He+ | 60 | 15 | 2276 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0848.log | He+ | 60 | 15 | 2355 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0845.log | He+ | 60 | 15 | 2465 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0843.log | He+ | 60 | 15 | 2550 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0841.log | He+ | 60 | 15 | 2649 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0830.log | He+ | 60 | 15 | 2661 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0832.log | He+ | 60 | 15 | 2753 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0834.log | He+ | 60 | 15 | 2845 | 7FE4 | 0.175 | -8 | 0 |
| Bern | PLA-FM2-050620-0839.log | He+ | 60 | 15 | 2995 | 7FE4 | 0.175 | -8 | 0 |

FM2:  Heavy Ions

Table of FM2 Heavy Ion Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bern | PLA-FM2-050624-0637.log | Ar5+ | 60 | 18 | 2171 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0640.log | Ar5+ | 60 | 18 | 2204 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0630.log | Ar5+ | 60 | 18 | 2242 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0525.log | Ar5+ | 60 | 18 | 2274 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0522.log | Ar5+ | 60 | 18 | 2353 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0519.log | Ar5+ | 60 | 18 | 2462 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0516.log | Ar5+ | 60 | 18 | 2553 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0514.log | Ar5+ | 60 | 18 | 2647 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0513.log | Ar5+ | 60 | 18 | 2750 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0510.log | Ar5+ | 60 | 18 | 2853 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0504.log | Ar5+ | 60 | 18 | 2990 | 7FE8 | 0.021 | -8 | 0 |
| Bern | PLA-FM2-050624-0924.log | Ne4+ | 60 | 20 | 2202 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0912.log | Ne4+ | 60 | 20 | 2272 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0908.log | Ne4+ | 60 | 20 | 2356 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0906.log | Ne4+ | 60 | 20 | 2461 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0904.log | Ne4+ | 60 | 20 | 2560 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0903.log | Ne4+ | 60 | 20 | 2642 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0901.log | Ne4+ | 60 | 20 | 2746 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0858.log | Ne4+ | 60 | 20 | 2840 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0855.log | Ne4+ | 60 | 20 | 3000 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-1008.log | Ne4+ | 60 | 15 | 2203 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-1001.log | Ne4+ | 60 | 15 | 2270 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0955.log | Ne4+ | 60 | 15 | 2348 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0952.log | Ne4+ | 60 | 15 | 2464 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0944.log | Ne4+ | 60 | 15 | 2550 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0942.log | Ne4+ | 60 | 15 | 2650 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0940.log | Ne4+ | 60 | 15 | 2758 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0939.log | Ne4+ | 60 | 15 | 2842 | 7FE8 | 0.026 | -8 | 0 |
| Bern | PLA-FM2-050624-0937.log | Ne4+ | 60 | 15 | 2986 | 7FE8 | 0.026 | -8 | 0 |

FM2 Solar Wind Sector Summary

To get an idea of the relative efficiencies for each species, they are plotted together in terms of energy per nucleon at the same MCP voltage setting (2650 V at Bern, which is equivalent to about 3100 V at UNH in Summer 2006).

FM2 Quadrants 1, 2, and 3

The other quadrants do not provide enough information to calculate start or single position efficiencies.  Below are plots for the stop efficiencies for quadrants 1, 2, and 3.

FM2 Q1 stop efficiency.  Data obtained from UNH calibration testing in Spring/Summer 2005.  The low rate for the 2 keV beams at -54 and -55 degrees are most likely the result of a physical blockage in the instrument near that location.

Table of FM2 Quadrant 1 Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UNH | PLA-FM2-050317-1849.log | air (Cu gun) | 2 | 15 | 2756 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050317-1848.log | air (Cu gun) | 2 | 15 | 2890 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050317-1846.log | air (Cu gun) | 2 | 15 | 3000 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050317-1823.log | air (Cu gun) | 2 | 15 | 3065 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050317-1831.log | air (Cu gun) | 2 | 15 | 3110 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050317-1834.log | air (Cu gun) | 2 | 15 | 3141 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050317-1836.log | air (Cu gun) | 2 | 15 | 3180 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050317-1844.log | air (Cu gun) | 2 | 15 | 3220 | 7F0E |  | -54 |  |
| UNH | PLA-FM2-050320-2204.log | air (45 gun) | 2 | 15 | 2955 | 7F0E |  | -55 |  |
| UNH | PLA-FM2-050320-2210.log | air (45 gun) | 2 | 15 | 2995 | 7F0E |  | -55 |  |
| UNH | PLA-FM2-050320-2214.log | air (45 gun) | 2 | 15 | 3062 | 7F0E |  | -55 |  |
| UNH | PLA-FM2-050320-2218.log | air (45 gun) | 2 | 15 | 3140 | 7F0E |  | -55 |  |
| UNH | PLA-FM2-050320-2220.log | air (45 gun) | 2 | 15 | 3198 | 7F0E |  | -55 |  |
| UNH | PLA-FM2-050323-2110.log | H+ (magnet gun) | 2 | 15 | 2265 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2107.log | H+ (magnet gun) | 2 | 15 | 2393 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2100.log | H+ (magnet gun) | 2 | 15 | 2460 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2058.log | H+ (magnet gun) | 2 | 15 | 2588 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2123.log | H+ (magnet gun) | 2 | 15 | 2735 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2057.log | H+ (magnet gun) | 2 | 15 | 2750 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2101.log | H+ (magnet gun) | 2 | 15 | 2845 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2049.log | H+ (magnet gun) | 2 | 15 | 2851 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2055.log | H+ (magnet gun) | 2 | 15 | 2851 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2103.log | H+ (magnet gun) | 2 | 15 | 2986 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2104.log | H+ (magnet gun) | 2 | 15 | 3048 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050323-2105.log | H+ (magnet gun) | 2 | 15 | 3171 | 7F0E |  | -76 |  |
| UNH | PLA-FM2-050324-1919.log | H+ (magnet gun) | 2 | 15 | 2262 | 7F0E |  | -111 |  |
| UNH | PLA-FM2-050324-1918.log | H+ (magnet gun) | 2 | 15 | 2457 | 7F0E |  | -111 |  |
| UNH | PLA-FM2-050324-1916.log | H+ (magnet gun) | 2 | 15 | 2754 | 7F0E |  | -111 |  |
| UNH | PLA-FM2-050324-1915.log | H+ (magnet gun) | 2 | 15 | 2819 | 7F0E |  | -111 |  |
| UNH | PLA-FM2-050324-1910.log | H+ (magnet gun) | 2 | 15 | 2990 | 7F0E |  | -111 |  |
| UNH | PLA-FM2-050324-1912.log | H+ (magnet gun) | 2 | 15 | 3046 | 7F0E |  | -111 |  |
| UNH | PLA-FM2-050607-0426.log | H+ (magnet gun) | 2 | 18 | 2319 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0430.log | H+ (magnet gun) | 2 | 18 | 2361 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0432.log | H+ (magnet gun) | 2 | 18 | 2446 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0434.log | H+ (magnet gun) | 2 | 18 | 2545 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0438.log | H+ (magnet gun) | 2 | 18 | 2631 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0440.log | H+ (magnet gun) | 2 | 18 | 2728 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0441.log | H+ (magnet gun) | 2 | 18 | 2816 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0443.log | H+ (magnet gun) | 2 | 18 | 2916 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-0445.log | H+ (magnet gun) | 2 | 18 | 2999 | 7F0E |  | -143 |  |
| UNH | PLA-FM2-050607-1856.log | H+ (magnet gun) | 30 | 18 | 2323 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1855.log | H+ (magnet gun) | 30 | 18 | 2368 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1854.log | H+ (magnet gun) | 30 | 18 | 2462 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1852.log | H+ (magnet gun) | 30 | 18 | 2572 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1850.log | H+ (magnet gun) | 30 | 18 | 2670 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1849.log | H+ (magnet gun) | 30 | 18 | 2757 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1847.log | H+ (magnet gun) | 30 | 18 | 2831 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1844.log | H+ (magnet gun) | 30 | 18 | 2928 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1831.log | H+ (magnet gun) | 30 | 18 | 3000 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1833.log | H+ (magnet gun) | 30 | 18 | 3051 | 7F76 |  | -151 |  |
| UNH | PLA-FM2-050607-1834.log | H+ (magnet gun) | 30 | 18 | 3106 | 7F76 |  | -151 |  |

Table of FM2 Quadrant 2 Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UNH | PLA-FM2-050607-1246.log | H+ (magnet gun) | 30 | 18 | 2072 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1245.log | H+ (magnet gun) | 30 | 18 | 2162 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1243.log | H+ (magnet gun) | 30 | 18 | 2252 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1241.log | H+ (magnet gun) | 30 | 18 | 2342 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1240.log | H+ (magnet gun) | 30 | 18 | 2431 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1237.log | H+ (magnet gun) | 30 | 18 | 2520 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1235.log | H+ (magnet gun) | 30 | 18 | 2611 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1234.log | H+ (magnet gun) | 30 | 18 | 2696 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1233.log | H+ (magnet gun) | 30 | 18 | 2790 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1231.log | H+ (magnet gun) | 30 | 18 | 2877 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1222.log | H+ (magnet gun) | 30 | 18 | 2965 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1228.log | H+ (magnet gun) | 30 | 18 | 3025 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-1230.log | H+ (magnet gun) | 30 | 18 | 3083 | 7F7A |  | -174 |  |
| UNH | PLA-FM2-050607-0330.log | H+ (magnet gun) | 2 | 18 | 2266 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0335.log | H+ (magnet gun) | 2 | 18 | 2360 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0337.log | H+ (magnet gun) | 2 | 18 | 2450 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0340.log | H+ (magnet gun) | 2 | 18 | 2549 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0343.log | H+ (magnet gun) | 2 | 18 | 2640 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0345.log | H+ (magnet gun) | 2 | 18 | 2727 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0347.log | H+ (magnet gun) | 2 | 18 | 2819 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0349.log | H+ (magnet gun) | 2 | 18 | 2907 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050607-0350.log | H+ (magnet gun) | 2 | 18 | 3000 | 7F0E |  | -185 |  |
| UNH | PLA-FM2-050320-2242.log | air (45 gun) | 2 | 15 | 2705 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050320-2241.log | air (45 gun) | 2 | 15 | 2812 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050320-2240.log | air (45 gun) | 2 | 15 | 2887 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050320-2237.log | air (45 gun) | 2 | 15 | 2951 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050320-2236.log | air (45 gun) | 2 | 15 | 2991 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050320-2230.log | air (45 gun) | 2 | 15 | 3041 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050320-2227.log | air (45 gun) | 2 | 15 | 3140 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050320-2223.log | air (45 gun) | 2 | 15 | 3198 | 7F0E |  | -152 |  |
| UNH | PLA-FM2-050323-2043.log | H+ (magnet gun) | 2 | 15 | 2456 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050323-2040.log | H+ (magnet gun) | 2 | 15 | 2592 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050323-2039.log | H+ (magnet gun) | 2 | 15 | 2750 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050323-2047.log | H+ (magnet gun) | 2 | 15 | 2851 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050323-2037.log | H+ (magnet gun) | 2 | 15 | 2991 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050323-2035.log | H+ (magnet gun) | 2 | 15 | 3046 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050323-2032.log | H+ (magnet gun) | 2 | 15 | 3176 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1808.log | H+ (magnet gun) | 2 | 15 | 2257 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1805.log | H+ (magnet gun) | 2 | 15 | 2458 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1804.log | H+ (magnet gun) | 2 | 15 | 2582 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1802.log | H+ (magnet gun) | 2 | 15 | 2744 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1800.log | H+ (magnet gun) | 2 | 15 | 2815 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1755.log | H+ (magnet gun) | 2 | 15 | 2981 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1757.log | H+ (magnet gun) | 2 | 15 | 3038 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050324-1759.log | H+ (magnet gun) | 2 | 15 | 3158 | 7F0E |  | -188 |  |
| UNH | PLA-FM2-050331-1951.log | air (big ion gun) | 31 | 18 | 2460 | 7F7A |  | -169 |  |
| UNH | PLA-FM2-050331-2000.log | air (big ion gun) | 31 | 18 | 2587 | 7F7A |  | -169 |  |
| UNH | PLA-FM2-050331-2007.log | air (big ion gun) | 31 | 18 | 2745 | 7F7A |  | -169 |  |
| UNH | PLA-FM2-050331-2010.log | air (big ion gun) | 31 | 18 | 2864 | 7F7A |  | -169 |  |
| UNH | PLA-FM2-050331-2013.log | air (big ion gun) | 31 | 18 | 3000 | 7F7A |  | -169 |  |
| UNH | PLA-FM2-050331-2015.log | air (big ion gun) | 31 | 18 | 3144 | 7F7A |  | -169 |  |
| UNH | PLA-FM2-050331-2019.log | air (big ion gun) | 31 | 18 | 3200 | 7F7A |  | -169 |  |

Table of FM2 Quadrant 3 Files

| Location | Filename | Species | Beam Voltage [keV] | PAC Voltage [kV] | MCP Voltage [V] | ESA Set [H'] | frequency/0.1cm^2 [kHz] average (single rate) | azimuth angle [degrees] | polar angle [degrees] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| UNH | PLA-FM2-050320-2252.log | air (Cu gun) | 2 | 15 | 2750 | 7F0E |  | 88 |  |
| UNH | PLA-FM2-050320-2245.log | air (Cu gun) | 2 | 15 | 2879 | 7F0E |  | 88 |  |
| UNH | PLA-FM2-050320-2254.log | air (Cu gun) | 2 | 15 | 2990 | 7F0E |  | 88 |  |
| UNH | PLA-FM2-050320-2256.log | air (Cu gun) | 2 | 15 | 3049 | 7F0E |  | 88 |  |
| UNH | PLA-FM2-050320-2259.log | air (Cu gun) | 2 | 15 | 3142 | 7F0E |  | 88 |  |
| UNH | PLA-FM2-050320-2301.log | air (Cu gun) | 2 | 15 | 3191 | 7F0E |  | 88 |  |
| UNH | PLA-FM2-050323-2013.log | H+ (magnet gun) | 2 | 15 | 2263 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-2017.log | H+ (magnet gun) | 2 | 15 | 2371 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-2007.log | H+ (magnet gun) | 2 | 15 | 2462 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-2005.log | H+ (magnet gun) | 2 | 15 | 2592 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-2003.log | H+ (magnet gun) | 2 | 15 | 2755 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-2000.log | H+ (magnet gun) | 2 | 15 | 2985 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-1952.log | H+ (magnet gun) | 2 | 15 | 3040 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-1957.log | H+ (magnet gun) | 2 | 15 | 3173 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050323-2024.log | H+ (magnet gun) | 2 | 15 | 3229 | 7F0E |  | 117 |  |
| UNH | PLA-FM2-050324-1839.log | H+ (magnet gun) | 2 | 15 | 2260 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050324-1836.log | H+ (magnet gun) | 2 | 15 | 2460 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050324-1834.log | H+ (magnet gun) | 2 | 15 | 2583 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050324-1833.log | H+ (magnet gun) | 2 | 15 | 2750 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050324-1831.log | H+ (magnet gun) | 2 | 15 | 2826 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050324-1852.log | H+ (magnet gun) | 2 | 15 | 2980 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050324-1828.log | H+ (magnet gun) | 2 | 15 | 3035 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050324-1830.log | H+ (magnet gun) | 2 | 15 | 3172 | 7F0E |  | 133 |  |
| UNH | PLA-FM2-050607-0328.log | H+ (magnet gun) | 2 | 18 | 2275 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0327.log | H+ (magnet gun) | 2 | 18 | 2357 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0325.log | H+ (magnet gun) | 2 | 18 | 2456 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0323.log | H+ (magnet gun) | 2 | 18 | 2549 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0320.log | H+ (magnet gun) | 2 | 18 | 2637 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0315.log | H+ (magnet gun) | 2 | 18 | 2729 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0313.log | H+ (magnet gun) | 2 | 18 | 2826 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0311.log | H+ (magnet gun) | 2 | 18 | 2909 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0304.log | H+ (magnet gun) | 2 | 18 | 3006 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050607-0310.log | H+ (magnet gun) | 2 | 18 | 3100 | 7F0E |  | 174 |  |
| UNH | PLA-FM2-050610-2109.log | H+ (magnet gun) | 2 | 18 | 2282 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2106.log | H+ (magnet gun) | 2 | 18 | 2467 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2105.log | H+ (magnet gun) | 2 | 18 | 2580 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2103.log | H+ (magnet gun) | 2 | 18 | 2761 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2059.log | H+ (magnet gun) | 2 | 18 | 2870 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2100.log | H+ (magnet gun) | 2 | 18 | 2948 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2102.log | H+ (magnet gun) | 2 | 18 | 3000 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2125.log | H+ (magnet gun) | 2 | 18 | 3008 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050610-2127.log | H+ (magnet gun) | 2 | 18 | 3057 | O6F8 |  | 170 |  |
| UNH | PLA-FM2-050607-1815.log | H+ (magnet gun) | 30 | 18 | 2170 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1813.log | H+ (magnet gun) | 30 | 18 | 2285 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1810.log | H+ (magnet gun) | 30 | 18 | 2376 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1808.log | H+ (magnet gun) | 30 | 18 | 2463 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1807.log | H+ (magnet gun) | 30 | 18 | 2572 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1805.log | H+ (magnet gun) | 30 | 18 | 2666 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1804.log | H+ (magnet gun) | 30 | 18 | 2759 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1803.log | H+ (magnet gun) | 30 | 18 | 2837 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1801.log | H+ (magnet gun) | 30 | 18 | 2921 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1758.log | H+ (magnet gun) | 30 | 18 | 3000 | 7F76 |  | 104 |  |
| UNH | PLA-FM2-050607-1759.log | H+ (magnet gun) | 30 | 18 | 3046 | 7F76 |  | 104 |  |